## BAM QC ST-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-06-19 |  | Model Number: | BAM-1020
TIME: | 10:45 |  | Serial Number: | P21275
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:46 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:45
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.61 | 0.5418% | -0.3599%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.1 | 23.2 | 0.9
Certification Expiration: | 2015-01-21 | 24.1 | 23.2 | 0.9
|  | 24.1 | 23.5 | 0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 747 | 749 | -2 | 628
Certification Expiration: | 2015-01-21 | 747 | 749 | -2 | 41276
|  | 747 | 749 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-06-19 |  | Model Number: | BAM-1020
TIME: | 10:55 |  | Serial Number: | P22904
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.3 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:56 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:57
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.86 | -0.949% | 1.14%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 22.7 | 23.2 | -0.5
Certification Expiration: | 2015-01-21 | 22.7 | 23.5 | -0.8
|  | 22.8 | 23.6 | -0.8
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 747 | 749 | -2 | 628
Certification Expiration: | 2015-01-21 | 747 | 749 | -2 | 41276
|  | 747 | 749 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-06-19 |  | Model Number: | BAM-1020
TIME: | 11:35 |  | Serial Number: | P22809
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 11:36 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 11:36
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.95 | -1.475% | 1.68%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.4 | 23.2 | 0.2
Certification Expiration: | 2015-01-21 | 23.4 | 23.2 | 0.2
|  | 23.4 | 23.2 | 0.2
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 746 | 749 | -3 | 628
Certification Expiration: | 2015-01-21 | 746 | 749 | -3 | 41276
|  | 746 | 749 | -3 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-N
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-N |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-06-19 |  | Model Number: | BAM-1020
TIME: | 10:40 |  | Serial Number: | P22906
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:41 |  |  | Certification Expiration: | 2015-01-21 | -100%
Reference Time: | 10:41
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.45 | 1.52% | -1.32%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 22 | 22.7 | -0.7
Certification Expiration: | 2015-01-21 | 22 | 22.7 | -0.7
|  | 22.2 | 22.8 | -0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 748 | 749 | -1 | 628
Certification Expiration: | 2015-01-21 | 748 | 749 | -1 | 41276
|  | 748 | 749 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-06-19 |  | Model Number: | BAM-1020
TIME: | 10:30 |  | Serial Number: | P22808
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.3 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:33 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:34
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.95 | -1.475% | 1.68%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 22.1 | 22.7 | -0.6
Certification Expiration: | 2015-01-21 | 22.1 | 22.7 | -0.6
|  | 22.1 | 22.7 | -0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 747 | 749 | -2 | 628
Certification Expiration: | 2015-01-21 | 747 | 749 | -2 | 41276
|  | 747 | 749 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM
DATE: | 2014-06-19 |  | Model Number: | BAM-1020
TIME: | 9:50 |  | Serial Number: | M4791
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 9:50 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 9:51
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.57 | 0.7846% | 1.037%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 19.3 | 20.5 | -1.2
Certification Expiration: | 2015-01-21 | 19.6 | 20.6 | -1
|  | 20.5 | 21.5 | -1
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 747 | 749 | -2 | 628
Certification Expiration: | 2015-01-21 | 747 | 749 | -2 | 41276
|  | 747 | 749 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Note: Changed time from 9:49 to 9:51.
Rev. 3 3/2014

## BAM QC NT-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-06-19 |  | Model Number: | 1020
TIME: | 9:20 |  | Serial Number: | M4793
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 9:21 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 9:21
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.57 | 0.7846% | -0.5999%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 18.3 | 18.8 | -0.5
Certification Expiration: | 2015-01-21 | 18.4 | 18.8 | -0.4
|  | 18.5 | 19.1 | -0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 748 | 749 | -1 | 628
Certification Expiration: | 2015-01-21 | 748 | 749 | -1 | 41276
|  | 748 | 749 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-06-19 |  | Model Number: | 1020
TIME: | 9:40 |  | Serial Number: | P21274
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.3 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 9:42 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 9:41
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.4 | 16.27 | 0.799% | -0.7927%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 19.6 | 19.7 | -0.1
Certification Expiration: | 2015-01-21 | 19.6 | 19.3 | 0.3
|  | 19.5 | 18.9 | 0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 747 | 749 | -2 | 628
Certification Expiration: | 2015-01-21 | 747 | 749 | -2 | 41276
|  | 747 | 749 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-06-19 |  | Model Number: | 1020
TIME: | 9:05 |  | Serial Number: | M4792
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 9:06 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 9:05
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.41 | 1.767% | 0.06098%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 19.4 | 19.3 | 0.1
Certification Expiration: | 2015-01-21 | 19.5 | 19.2 | 0.3
|  | 19.1 | 18.5 | 0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 749 | 0 | 628
Certification Expiration: | 2015-01-21 | 749 | 749 | 0 | 41276
|  | 749 | 749 | 0 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014