## ST-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Metals
DATE: | 2014-04-25 |  | Model Number: | 2025i
TIME: | 9:40 |  | Serial Number: | 20547
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 14.6 | 15 | -0.4
Certification Expiration: | 2015-01-21 | 14.3 | 15 | -0.7
|  | 14.3 | 15.1 | -0.8
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 739 | 740 | -1 | 637
Certification Expiration: | 2015-01-21 | 739 | 740 | -1 | 41285
|  | 739 | 740 | -1 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 13.8 | 13.5 | 0.3
Certification Expiration: | 2015-01-21 | 13.8 | 13.5 | 0.3
|  | 13.8 | 13.5 | 0.3
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 9 | mmHg | Partisol : | 10:06
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 10:06
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.67 | 17.14 | -2.742% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Note: Initial leak check reading was 26. Cleaned all seals and reading dropped 5. |  |  |  |  | Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## ST-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Carbon
DATE: | 2014-04-25 |  | Model Number: | 2025i
TIME: | 9:10 |  | Serial Number: | 20544
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 13.8 | 12.8 | 1
Certification Expiration: | 2015-01-21 | 13.9 | 12.8 | 1.1
|  | 13.8 | 12.7 | 1.1
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 741 | 740 | 1 | 637
Certification Expiration: | 2015-01-21 | 741 | 740 | 1 | 41285
|  | 741 | 740 | 1 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 13 | 12.8 | 0.2
Certification Expiration: | 2015-01-21 | 13 | 12.8 | 0.2
|  | 13.1 | 12.8 | 0.3
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 18 | mmHg | Partisol : | 9:22
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 9:22
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.69 | 16.51 | 1.09% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Note: Changed time on sampler from 09:20 to 09:22.
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Metals
DATE: | 2014-04-15 |  | Model Number: | 2025i
TIME: | 10:30 |  | Serial Number: | 20546
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 16.7 | 17.2 | -0.5
Certification Expiration: | 2015-01-21 | 16.7 | 17.2 | -0.5
|  | 16.6 | 17.1 | -0.5
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 740 | 741 | -1 | 636
Certification Expiration: | 2015-01-21 | 740 | 741 | -1 | 41284
|  | 740 | 741 | -1 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 17.5 | 18.6 | -1.1
Certification Expiration: | 2015-01-21 | 17.5 | 18.6 | -1.1
|  | 17.5 | 18.6 | -1.1
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 15 | mmHg | Partisol : | 10:38
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 10:38
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.68 | 16.82 | -0.8323% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Note: Changed time on sampler from 09:36 to 09:38. |  |  |  |  | Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Carbon
DATE: | 2014-04-25 |  | Model Number: | 2025i
TIME: | 10:45 |  | Serial Number: | 20545
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 16.2 | 16.6 | -0.4
Certification Expiration: | 2015-01-21 | 16.3 | 16.6 | -0.3
|  | 16.2 | 16.5 | -0.3
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 740 | 742 | -2 | 635
Certification Expiration: | 2015-01-21 | 740 | 742 | -2 | 41283
|  | 740 | 742 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 17.7 | 17.2 | 0.5
Certification Expiration: | 2015-01-21 | 17.7 | 17.2 | 0.5
|  | 17.7 | 17.2 | 0.5
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
|  |  | Clock/Timer Verification5: | 11:37
External Leak Check4 | 14 | mmHg | Partisol : | 10:48
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 10:48
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.68 | 16.53 | 0.9074% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM