## ST-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Metals
DATE: | 2014-06-24 |  | Model Number: | 2025i
TIME: | 11:15 |  | Serial Number: | 20547
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 29 | 28.6 | 0.4
Certification Expiration: | 2015-01-21 | 29.2 | 28.6 | 0.6
|  | 29.2 | 28.5 | 0.7
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 741 | 744 | -3 | 633
Certification Expiration: | 2015-01-21 | 741 | 744 | -3 | 41281
|  | 741 | 744 | -3 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 28.8 | 29 | -0.2
Certification Expiration: | 2015-01-21 | 29.1 | 29.1 | 0
|  | 29.2 | 29.2 | 0
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 12 | mmHg | Partisol : | 11:18
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 11:20
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 15.12 | 15.16 | -0.2639% | 0%
|  | 18.39 | 18.71 | -1.71% | 0%
|  | 16.66 | 16.84 | -1.069% | 0%
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Note: Time changed from 11:18 to 11:20. |  |  |  |  | Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## ST-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Carbon
DATE: | 2014-06-24 |  | Model Number: | 2025i
TIME: | 10:30 |  | Serial Number: | 20544
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 29.3 | 29.8 | -0.5
Certification Expiration: | 2015-01-21 | 29.3 | 29.8 | -0.5
|  | 29.3 | 29.7 | -0.4
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 744 | 744 | 0 | 633
Certification Expiration: | 2015-01-21 | 744 | 744 | 0 | 41281
|  | 744 | 744 | 0 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 28.7 | 29.2 | -0.5
Certification Expiration: | 2015-01-21 | 28.7 | 29.3 | -0.6
|  | 28.7 | 29.2 | -0.5
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 23 | mmHg | Partisol : | 10:31
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 10:31
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 15.11 | 14.63 | 3.281% | 0%
|  | 18.41 | 18.35 | 0.327% | 0%
|  | 16.69 | 16.31 | 2.33% | 0%
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Metals
DATE: | 2014-06-24 |  | Model Number: | 2025i
TIME: | 8:00 |  | Serial Number: | 20546
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23 | 23.2 | -0.2
Certification Expiration: | 2015-01-21 | 23 | 23.1 | -0.1
|  | 23 | 22.9 | 0.1
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 741 | 743 | -2 | 634
Certification Expiration: | 2015-01-21 | 741 | 743 | -2 | 41282
|  | 741 | 743 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.7 | 24.2 | -0.5
Certification Expiration: | 2015-01-21 | 23.8 | 24.2 | -0.4
|  | 23.8 | 24.2 | -0.4
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 20 | mmHg | Partisol : | 8:15
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 8:15
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 15.09 | 14.7 | 2.653% | 0%
|  | 18.41 | 17.98 | 2.392% | 0%
|  | 16.68 | 16.79 | -0.6552% | 0%
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Carbon
DATE: | 2014-06-24 |  | Model Number: | 2025i
TIME: | 8:45 |  | Serial Number: | 20545
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.9 | 23.8 | 0.1
Certification Expiration: | 2015-01-21 | 23.7 | 24 | -0.3
|  | 23.8 | 24 | -0.2
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 741 | 743 | -2 | 634
Certification Expiration: | 2015-01-21 | 741 | 743 | -2 | 41282
|  | 741 | 743 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 25 | 24.7 | 0.3
Certification Expiration: | 2015-01-21 | 25.1 | 24.7 | 0.4
|  | 25 | 24.7 | 0.3
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
|  |  | Clock/Timer Verification5: | 11:37
External Leak Check4 | 21 | mmHg | Partisol : | 9:00
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 9:00
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 15.1 | 14.71 | 2.651% | 0%
|  | 18.39 | 18.01 | 2.11% | 0%
|  | 16.67 | 16.28 | 2.396% | 0%
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM