## ST-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Metals
DATE: | 2014-07-18 |  | Model Number: | 2025i
TIME: | 8:30 |  | Serial Number: | 20547
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 26.6 | 26.7 | -0.1
Certification Expiration: | 2015-01-21 | 26.6 | 26.5 | 0.1
|  | 26.5 | 26.2 | 0.3
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 751 | -2 | 626
Certification Expiration: | 2015-01-21 | 749 | 751 | -2 | 41274
|  | 749 | 751 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 26.6 | 26.7 | -0.1
Certification Expiration: | 2015-01-21 | 26.7 | 27.1 | -0.4
|  | 26.7 | 27.2 | -0.5
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 19 | mmHg | Partisol : | 8:45
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 8:45
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.68 | 16.89 | -1.243% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## ST-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Carbon
DATE: | 2014-07-18 |  | Model Number: | 2025i
TIME: | 10:10 |  | Serial Number: | 20544
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 26.1 | 25.3 | 0.8
Certification Expiration: | 2015-01-21 | 26.1 | 25.3 | 0.8
|  | 26.2 | 25.5 | 0.7
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 751 | 751 | 0 | 626
Certification Expiration: | 2015-01-21 | 751 | 751 | 0 | 41274
|  | 751 | 751 | 0 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 26.8 | 26.8 | 0
Certification Expiration: | 2015-01-21 | 26.8 | 27.1 | -0.3
|  | 26.8 | 27.2 | -0.4
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 21 | mmHg | Partisol : | 10:32
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 10:32
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.67 | 16.35 | 1.957% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Metals
DATE: | 2014-07-18 |  | Model Number: | 2025i
TIME: | 11:20 |  | Serial Number: | 20546
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.3 | 24.8 | -0.5
Certification Expiration: | 2015-01-21 | 24.4 | 24.7 | -0.3
|  | 24.1 | 24.5 | -0.4
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 748 | 751 | -3 | 626
Certification Expiration: | 2015-01-21 | 748 | 751 | -3 | 41274
|  | 748 | 751 | -3 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 27.1 | 26.2 | 0.9
Certification Expiration: | 2015-01-21 | 27.1 | 26.5 | 0.6
|  | 27.2 | 26.5 | 0.7
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 20 | mmHg | Partisol : | 11:25
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 11:25
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.67 | 17.08 | -2.4% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Carbon
DATE: | 2014-07-18 |  | Model Number: | 2025i
TIME: | 11:50 |  | Serial Number: | 20545
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.2 | 25.1 | -0.9
Certification Expiration: | 2015-01-21 | 24.3 | 25 | -0.7
|  | 24 | 24.8 | -0.8
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 751 | -2 | 626
Certification Expiration: | 2015-01-21 | 749 | 751 | -2 | 41274
|  | 749 | 751 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 27.1 | 26.7 | 0.4
Certification Expiration: | 2015-01-21 | 27.1 | 26.7 | 0.4
|  | 27 | 26.8 | 0.2
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 9 | mmHg | Partisol : | 11:51
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 11:52
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.67 | 16.51 | 0.9691% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Note: Changed time from 11:51 to 11:52 CST.
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM