## VOA
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2027
FRACTION: | VOA (SW846 8260C) for the Target Compound List only
MEDIA: | GROUNDWATER
SAMPLES: | MW-103, PZ-200-SS, S-61, PZ-115-SS, MW-104, PZ-100-SD, I-67, PZ-203-SS, PZ-100-SS, I-66, PZ-305-AI, DUP 03, DUP 04
SAMPLE DATE(S): | 2013-04-05
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times, Preservation | Cooler temp < 6 ºC. Samples holding time requirement < 7 days (<14 days if HCL preserved). | Coolers received at 3° C. All holding time criteria met. | None
System Monitoring Compounds | recoveries within limits (70 - 130%) or laboratory established limits | System monitoring compound recoveries were in control. | None
Matrix Spike/Matrix Spike Duplicates | MS/MSD: 1 per 20 project samples. Recoveries and RPDs within lab limits | MS/MSD performed on MW-103 and I-66 - all results in control for target compounds | None
Lab Control Sample/Duplicate | Recoveries within lab limits (or 70-130%). | LCS results for target compounds were within control limits | None
Blanks | Method blanks: 1 per 12 hours. No analytes detected in MB, TB, or EB. | MB 160-45293/2 all compounds were not detected MB 160-46058/3-A all compounds were not detected | None
| Trip Blanks | Ethylbenzene 0.31 ug/L - qualify associated samples as U at RL | Ethylbenzene 5 U MW-103, PZ-200-SS, S-61, PZ-115-SS, PZ-100-SD, PZ-203-SS, PZ-100-SS, I-66, PZ-305-AI, DUP 03, DUP 04
| Equipment Blanks | No equipment blanks were associated to these samples. | None
GC/MS Initial Calibration | %RSD ≤ 30%. Average RRFs > 0.050. | Initial Calibration Instrument MSF 3/25/13: All %RSD criteria met. Average RRF meets requirements. | None
GC/MS Continuing Calibration | CV performed for every 12 hours per instrument. %D ≤ 25%. RRFs ≥ 0.05. | CCVIS 160-45293/3 - results were in control CCVIS 160-46058/2 - results were in control RF was greater than or equal to 0.05 ( or 0.01 for poor response compounds) for all target compounds | None
Internal Standards | IS areas of samples & blank within (-50% to + 100%). RTs < 30 seconds. | Standard recovery area within the QC limits for all three standards; and retention times were within 30 seconds of the standard for all samples that were used in this SDG. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP 03 was collected as a field dup of I-67 - all results were in control DUP 04 was collected as a field dup of PZ-305-AI- all results were in control | None
RT = Retention Time; %D = Percent Deviation; %RPD = Relative Percent Difference; %RSD = Percent Relative Standard Deviation; RRF = Relative Response Factor; CCV = Continuing Calibration Verification
TCL = Target Compound List; TIC = Tentatively Identified Compound

## Metals
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2027
FRACTION: | Metals (SW846 6010B and 7470A/7471A)
MEDIA: | GROUNDWATER (TOTAL)
SAMPLES: | MW-103, PZ-200-SS, S-61, PZ-115-SS, MW-104, PZ-100-SD, I-67, PZ-203-SS, PZ-100-SS, I-66, PZ-305-AI, DUP 03, DUP 04
SAMPLE DATE(S): | 2013-04-05
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times & Preservation | Cooler temp < 4 C. Holding Time: Hg < 28 days, all other metals < 180 days from collection. | Coolers received at 3° C. All holding time criteria met. | None
Calibration | Calibration every 24 hours Hg r^2 > 0.995 CCV every 10 samples or 2 hours ICP CV/CCV %R = 90-110% Hg CV/CCV %R = 80-120% | Hg calibrations, ICVs, and CCVs were in control Metals ICVs and CCVs were in control | None
Blanks (method blank, prep blank) | CC Blank Conc < CRQL. Method blanks: 1 per 20 project samples. No TCL detected in MB, TB, or EB. | Method blanks: MB 160-45567/1-A (total) - all analytes were not detected MB 160-45569/1-A (dissolved) Sb 4.10 ug/L - qualify associated sample detects < RL as U Fe 41.4 ug/L - qualify associated sample detects < RL as U Zn 6.8 ug/L - qualify associated sample detects < RL as U MB 160-46413/1-A: Hg (Total) - not detected MB 160-46718/1-A: Hg (dissolved) - not detected Equipment blank: No equipment blanks associated to this SDG. CCBs: CCB 45942/73 (tot) Sb (6.7 ug/L) - associated samples were not detected Be (0.8 ug/L) - associated samples were not detected Se (7.2 ug/L) - qualify associated sample detects < RL as U | MW-103 Sb (dis) 10U Zn (dis) 20U Se (tot) 75 U As (dis) 10U Se (dis) 15 U V (dis) 50U PZ-200-SS Sb (dis) 10U Zn (dis) 20U Se (dis) 15 U S-61 Zn (dis) 20U PZ-115-SS Zn (dis) 20U As (dis) 10U MW-104 Sb (dis) 10U Zn (dis) 20U Se (dis) 15 U
|  | CCB 45942/85 (tot) Sb (6.3 ug/L) - associated samples were not detected Be (1.1 ug/L) - associated samples were not detected Se (5.2 ug/L) - qualify associated sample detects < RL as U CCB 45942/97 (tot) Sb (6.7 ug/L) - associated samples were not detected Be (0.7 ug/L) - associated samples were not detected Se (7.6 ug/L) - associated samples were not detected CCB 46208/13 (dis) Sb (7.9 ug/L) - qualify associated sample detects < RL as U As (4.1 ug/L) - qualify associated sample detects < RL as U Be (1.1 ug/L) - associated samples were not detected Cd (1.0 ug/L) - associated samples were not detected Se (7.8 ug/L) - qualify associated sample detects < RL as U V (5.5 ug/L) - qualify associated sample detects < RL as U CCB 46208/25 (dis) Sb (4.1 ug/L)- qualify associated sample detects < RL as U As (2.4 ug/L) - qualify associated sample detects < RL as U Se (6.4 ug/L) - qualify associated sample detects < RL as U | PZ-100-SD Zn (dis) 20U As (dis) 10U I-67 Zn (dis) 20U PZ-203-SS Sb (dis) 10U Zn (dis) 20U PZ-100-SS Zn (dis) 20U I-66 Zn (dis) 20U As (dis) 10U Se (dis) 15 U PZ-305-AI Sb (dis) 10U Zn (dis) 20U As (dis) 14 J+ V (dis) 50U
|  | CCB 46208/37 (dis) Be (0.8 ug/L) - associated samples were not detected As (2.9 ug/L) - qualify associated sample detects < RL as U Se (5.7 ug/L) - qualify associated sample detects < RL as U CCB 46208/49 (dis) Sb (5.2 ug/L) - qualify associated sample detects < RL as U As (2.0 ug/L) - qualify associated sample detects < RL as U Be (1.0 ug/L) - associated samples were not detected Se (7.6 ug/L) - qualify associated sample detects < RL as U V (5.2 ug/L) - associated samples were not detected | DUP 03 As (dis) 10U DUP 04 Sb (dis) 10U Zn (dis) 20U V (dis) 50U
CRDL Standard | CRDL results btw 70-130% 50-150% for Sb, Pb, and Tl | The CRI was in control | None
Laboratory Control Sample | LCS/LCSD: 1 per 20 project samples or each preparation batch. LCS limits within 80-120%. | LCS results within limits for total and dissolved metals | None
Lab Duplicates | RPD <20% or Absolute Diff < RL when samp/dup value < 5x RL | No lab duplicates were analyzed within this SDG. | None
Matrix Spike/Matrix Spike Duplicates | MS/MSD: 1 per 20 project samples or each preparation batch. Recoveries within lab limits. MS/MSD %RPDs <= 20%. Post Digestion Spike Recovery (when MS spike rec out of control) limits 75-125% | MS/MSD analysis for total 6010 metals performed on MW-103 Al 158%/145% - qualify associated sample conc as J+ K 184%/181% - sample concentration was not detected - no qualification required - Some analytes not calculated because the sample conc was > 4X the spike conc - apply qualifications to parent sample only (entire batch not qualified because another spike sample was in control) MS/MSD analysis for dissolved 6010 metals performed on MW-103 - results were in control Some analytes not calculated because the sample conc was > 4X the spike conc MS/MSD analysis for total mercury performed on PZ-200-SS- results were in control MS/MSD analysis for dissolved mercury performed on S-61 - results were in control | MW-103 Al (tot) 13000 J+
ICP Interference Check Sample (ICS) | ICS results within 80-120%. | All criteria met. | None
Linear Range Check Standard | results within 90-110% | Linear range check standards were in control. | None
Serial Dilution | Performed on samples of a similar matrix or 1 per 20 samples. %D ≤ 10% conc ≥ 50xDL for 5-fold dilution. | Serial dilutions were performed on MW-103 and PZ-200-SS (Hg). Results were in control. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP 03 was collected as a field dup of I-67 - all results were in control DUP 04 was collected as a field dup of PZ-305-AI- Al (tot) RPD = 40% - qualify parent sample and DUP as J | PZ-305-AI Al (tot) 6500 J DUP 04 Al (tot) 9800 J
RT = Retention Time; %D = Percent Deviation; %RPD = Relative Percent Difference; %RSD = Percent Relative Standard Deviation; RRF = Relative Response Factor; CCV = Continuing Calibration Verification
TCL = Target Compound List; MS = Matrix Spike; MSD = Matrix Spike Duplicate;

## WetChem
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2027
FRACTION: | Chloride, Sulfate, Nitrate, Bromide and Iodide (EPA 300)/ Alkalinity (EPA 310.1)
MEDIA: | GROUNDWATER
SAMPLES: | MW-103, PZ-200-SS, S-61, PZ-115-SS, MW-104, PZ-100-SD, I-67, PZ-203-SS, PZ-100-SS, I-66, PZ-305-AI, DUP 03, DUP 04
SAMPLE DATE(S): | 2013-04-05
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times & Preservation | Cooler temp < 6 C. | Coolers received at 3° C. Holding time exceeded for nitrate on some samples Holding time requirements were not met for all samples. | MW-103 NO3 0.053 J PZ-200-SS NO3 ND R S-61 NO3 ND R PZ-115-SS NO3 0.0076 J MW-104 NO3 0.019 J PZ-100-SD NO3 0.017 J I-67 NO3 0.0062 J PZ-203-SS NO3 ND R PZ-100-SS NO3 0.0088 J I-66 NO3 ND R PZ-305-AI NO3 ND R DUP 03 NO3 0.16 J DUP 04 NO3 0.0059 J
Holding time = 28 days for chloride, sulfate, bromide and iodide 14 days for alkalinity 48 hrs for nitrate.
Calibration and Continuing Calibration Verification Standards | R2>0.99 CCVs 90-110% | ICVs, CCVs, and Calibrations were in control. | None
Blanks | Method blanks: 1 per 20 project samples. No analytes detected in MBs, CCBs or EBs. | Method Blank: MB 160-44533/1 ALK 0.250 mg/L - sample concs > 5X blank conc X DF- no qualification required. Equipment blank: No equipment blanks were collected for these methods. CCBs: CCB 160/45380/91 Cl 0.0738 mg/L - sample concs > 5X blank conc X DF- no qualification required. | None
Laboratory Control Sample | Recoveries within lab limits. | LCS recoveries were in control. | None
| Lab Duplicate: RPD <20% or Absolute Diff < RL when samp/dup value < 5x RL | MW-104 (NO3, I, Cl, SO4, ALK & Br) were used as lab duplicates. Results were in control. | None
Lab Duplicates
Spike Sample Analysis | MS recovery within lab established limits | MS analysis for bromide, ALK, Cl, SO4 and NO3 was performed on MW-104 - results were in control MS analysis for Iodide was performed on PZ-203-SS - results were in control MS analysis for Cl was performed on PZ-305-AI - results were in control | None
Low Level Control Sample | recovery within 80-120% | Performed on alkalinity only. Results were in control. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP 03 was collected as a field dup of I-67 - NO3 = %185 - qualify parent sample and DUP as J DUP 04 was collected as a field dup of PZ-305-AI- - all results were in control | DUP 03 NO3 0.16 J I-67 NO3 0.0062 J

## field dups
chemical_name | DUP 03 |  | I-67 |  | range | rpd |  | DUP 04 |  | PZ-305-AI |  | range | rpd
ALKALINITY, TOTAL (AS CACO3) | 610 | 5 | 620 | 5 | -10 | -1.626% | ALKALINITY, TOTAL (AS CACO3) | 950 | 5 | 980 | 5 | -30 | -3.109%
ALUMINUM | 820 | 1000 | 950 | 1000 | -130 | -14.69% | ALUMINUM | 9800 | 1000 | 6500 | 1000 | 3300 | 40.49%
ARSENIC | 2.7 | 10 |  |  | 2.7 |  | ANTIMONY | 5.4 | 10 | 4.3 | 10 | 1.1 | 22.68%
BARIUM | 260 | 250 | 250 | 250 | 10 | 3.922% | ARSENIC | 16 | 10 | 14 | 10 | 2 | 13.33%
BROMIDE | 0.097 | 0.25 | 0.11 | 0.25 | -0.013 | -12.56% | ARSENIC | 17 | 50 | 11 | 50 | 6
CALCIUM | 240000 | 5000 | 240000 | 5000 | 0 | 0% | BARIUM | 930 | 250 | 700 | 50 | 230 | 28.22%
ETHYLBENZENE | 0.38 | 5 |  |  | 0.38 |  | BARIUM | 690 | 50 | 820 | 250 | -130 | -17.22%
IRON | 4800 | 100 | 4400 | 100 | 400 | 8.696% | BENZENE | 1.2 | 5 | 1.2 | 5 | 0 | 0%
LEAD | 2.5 | 10 | 2.4 | 10 | 0.1 | 4.082% | BROMIDE | 0.58 | 0.25 | 0.6 | 0.25 | -0.02 | -3.39%
MAGNESIUM | 35000 | 5000 | 35000 | 5000 | 0 | 0% | CALCIUM | 290000 | 10000 | 290000 | 10000 | 0 | 0%
MANGANESE | 1200 | 15 | 1200 | 15 | 0 | 0% | CALCIUM | 300000 | 10000 | 290000 | 10000 | 10000 | 3.39%
MANGANESE | 1200 | 75 | 1200 | 75 | 0 | 0% | CHLOROBENZENE | 3.9 | 5 | 3.9 | 5 | 0 | 0%
NITROGEN, NITRATE (AS N) | 0.16 | 0.02 | 0.0062 | 0.02 | 0.1538 | 185.1% | CHROMIUM, TOTAL | 23 | 50 |  |  | 23
Potassium | 9900 | 5000 | 10000 | 5000 | -100 | -1.005% | COBALT | 4.1 | 50 |  |  | 4.1
SODIUM | 63000 | 1000 | 64000 | 1000 | -1000 | -1.575% | COPPER | 23 | 130 |  |  | 23
|  |  | 11 | 20 | -11 |  | ETHYLBENZENE | 0.31 | 5 | 0.41 | 5 | -0.1 | -27.78%
SULFATE (AS SO4) | 62 | 10 | 63 | 10 | -1 | -1.6% | IODIDE (AS I) | 0.41 | 1 | 0.37 | 1 | 0.04 | 10.26%
CHLORIDE (AS CL) | 130 | 20 | 130 | 20 | 0 | 0% | IRON | 45000 | 500 | 42000 | 500 | 3000 | 6.897%
[7 empty] | IRON | 34000 | 100 | 34000 | 100 | 0 | 0%
[7 empty] | LEAD | 5.2 | 10 | 4.8 | 10 | 0.4 | 8%
[7 empty] | LEAD | 27 | 50 | 19 | 50 | 8
[7 empty] | MAGNESIUM | 80000 | 5000 | 78000 | 10000 | 2000 | 2.532%
[7 empty] | MAGNESIUM | 77000 | 10000 | 78000 | 5000 | -1000 | -1.29%
[7 empty] | MANGANESE | 3100 | 15 | 3200 | 75 | -100 | -3.175%
[7 empty] | MANGANESE | 3300 | 75 | 3100 | 15 | 200 | 6.25%
[7 empty] | NITROGEN, NITRATE (AS N) | 0.0059 | 0.02 |  |  | 0.0059
[7 empty] | Potassium | 10000 | 25000 | 9100 | 25000 | 900 | 9.424%
[7 empty] | Potassium | 7200 | 5000 | 7100 | 5000 | 100 | 1.399%
[7 empty] | SODIUM | 19000 | 1000 | 19000 | 1000 | 0 | 0%
[7 empty] | SODIUM | 19000 | 5000 | 19000 | 5000 | 0 | 0%
[7 empty] | SULFATE (AS SO4) | 1.8 | 0.5 | 1.8 | 0.5 | 0 | 0%
[7 empty] | VANADIUM | 4.3 | 50 | 5.6 | 50 | -1.3 | -26.26%
[7 empty] | ZINC | 8.7 | 20 | 12 | 100 | -3.3 | -31.88%
[7 empty] | CHLORIDE (AS CL) | 65 | 4 | 66 | 4 | -1 | -1.527%