## VOA
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2044
FRACTION: | VOA (SW846 8260C) for the Target Compound List only
MEDIA: | GROUNDWATER
SAMPLES: | PZ-103-SS, PZ-114-AS, PZ-201A-SS, PZ-204A-SS, PZ-205-AS, PZ-205-SS, PZ-206-SS, PZ-204-SS, I-68, D-87, PZ-106-SD, S-82, PZ-106-SS, I-9, D-93, PZ-111-KS, D-6, D-83, DUP05
SAMPLE DATE(S): | 4/8,9/2013
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times, Preservation | Cooler temp < 6 ºC. Samples holding time requirement < 7 days (<14 days if HCL preserved). | Coolers received at 2° C. All holding time criteria met. | None
System Monitoring Compounds | recoveries within limits (70 - 130%) or laboratory established limits | System monitoring compound recoveries were in control. Surrogate recoveries failed in the first analysis of I-68, I-9, and PZ-205-SS. The samples were reextracted and reanalyzed and the reanalyzed results reported. | None
Matrix Spike/Matrix Spike Duplicates | MS/MSD: 1 per 20 project samples. Recoveries and RPDs within lab limits | MS/MSD performed on: DUP05 Bromomethane MS = 67%- qualify associated sample as UJ- 2-Hexanone MSD = 122% - no qualification required because result was ND Acetone RPD = 34% - no qualification required because result was ND Chloroethane MS = 72%- qualify associated sample as UJ- 1,2-Dibromo-3-chloropropane RPD = 21% - no qualification required because result was ND PZ-204-SS - all results in control for target compounds PZ-205-SS - all results in control for target compounds Acetone RPD = 33% - no qualification required because result was ND | DUP05 Bromomethane UJ- Chloroethane UJ-
Lab Control Sample/Duplicate | Recoveries within lab limits (or 70-130%). | LCS results for target compounds were within control limits except for: LCS 160-45727/4 Ethylbenzene 84% - qualify associated samples as J-/UJ- LCS 160-46184/4 1,1-Dichloroethene 82%- qualify associated samples as J-/UJ- Bromomethane 67%- qualify associated samples as J-/UJ- | PZ-103-SS, PZ-114-AS, PZ-201A-SS, PZ-204A-SS, PZ-205-AS, PZ-206-SS, PZ-204-SS, D-87, PZ-106-SD, S-82, PZ-106-SS, D-93, PZ-111-KS, D-6, D-83, DUP05 Ethylbenzene J-/UJ- I-68, I-9, and PZ-205-SS 1,1-Dichloroethene J-/UJ- Bromomethane J-/UJ-
Blanks | Method blanks: 1 per 12 hours. No analytes detected in MB, TB, or EB. | MB 160-45595/2 all compounds were not detected MB 160-45727/2 all compounds were not detected MB 160-46184/2 all compounds were not detected | None
| Trip Blanks | all compounds were not detected | None
| Equipment Blanks | FB at PZ-201A-SS - all compounds were not detected | None
GC/MS Initial Calibration | %RSD ≤ 30%. Average RRFs > 0.050. | Initial Calibration Instrument MSF 3/25/13: All %RSD criteria met. Average RRF meets requirements. Instrument MSX 3/25/13: All %RSD criteria met. Average RRF meets requirements. | None
GC/MS Continuing Calibration | CV performed for every 12 hours per instrument. %D ≤ 25%. RRFs ≥ 0.05. | CCVIS 160-45595/3 - results were in control CCVIS 160-46184/3 - results were in control CCVIS 160-45727/3 - results were in control RF was greater than or equal to 0.05 ( or 0.01 for poor response compounds) for all target compounds | None
Internal Standards | IS areas of samples & blank within (-50% to + 100%). RTs < 30 seconds. | Standard recovery area within the QC limits for all three standards; and retention times were within 30 seconds of the standard for all samples that were used in this SDG. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP05 was collected as a field dup of I-9 - all results were in control | None
RT = Retention Time; %D = Percent Deviation; %RPD = Relative Percent Difference; %RSD = Percent Relative Standard Deviation; RRF = Relative Response Factor; CCV = Continuing Calibration Verification
TCL = Target Compound List; TIC = Tentatively Identified Compound

## Metals
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2044
FRACTION: | Metals (SW846 6010B and 7470A/7471A)
MEDIA: | GROUNDWATER (TOTAL)
SAMPLES: | PZ-103-SS, PZ-114-AS, PZ-201A-SS, PZ-204A-SS, PZ-205-AS, PZ-205-SS, PZ-206-SS, PZ-204-SS, I-68, D-87, PZ-106-SD, S-82, PZ-106-SS, I-9, D-93, PZ-111-KS, D-6, D-83, DUP05
SAMPLE DATE(S): | 4/8,9/2013
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times & Preservation | Cooler temp < 4 C. Holding Time: Hg < 28 days, all other metals < 180 days from collection. | Coolers received at 2° C. All holding time criteria met. S-82 was received at the laboratory without preservative. The laboratory corrected the pH to <2. For a reason the lab didn't explain, some of the dissolved metals were analyzed 3 times. Only the final results were validated. | None
Calibration | Calibration every 24 hours Hg r^2 > 0.995 CCV every 10 samples or 2 hours ICP CV/CCV %R = 90-110% Hg CV/CCV %R = 80-120% | Hg calibrations, ICVs, and CCVs were in control Metals ICVs and CCVs were in control | None
Blanks (method blank, prep blank) | CC Blank Conc < CRQL. Method blanks: 1 per 20 project samples. No TCL detected in MB, TB, or EB. | Method blanks: MB 160-45572/1-A (total) Zn 6.3 ug/L - qualify associated sample detects < RL as U - qualify associated sample detects> RL but <5X blank conc X DF as J+ MB 160-45573/1-A (dissolved) - no analytes detected MB 160-46719/1-A: Hg (Total) - not detected MB 160-47100/1-A: Hg (dissolved) - not detected Equipment blank: No equipment blanks associated to this SDG. CCBs: CCB 46208/37 (tot) As (2.9 ug/L) - qualify associated sample detects < RL as U Be (0.8 ug/L) - qualify associated sample detects < RL as U Se (5.7 ug/L) - - associated samples were not detected | PZ-103-SS Zn (tot)110 J+ As (tot) 50 U V (tot) 250 U Se (dis) 75U PZ-114-AS Zn (tot) 100 U V (tot) 250 U PZ-201-AS Zn (tot) 100 U As (tot) 50 U PZ-204A-SS Zn (tot) 100 U As (tot) 50 U Be (tot) 25 U
|  | CCB 46208/49 (tot) Sb (5.2 ug/L)- qualify associated sample detects < RL as U As (2.0 ug/L) - qualify associated sample detects < RL as U Be (1.0 ug/L) -qualify associated sample detects < RL as U Se (7.6 ug/L) - associated samples were not detected V (5.2 ug/L) - qualify associated sample detects < RL as U CCB 46208/61 (tot) Sb (4.5 ug/L)- qualify associated sample detects < RL as U As (2.7 ug/L) - qualify associated sample detects < RL as U Be (0.8 ug/L) - associated samples were not detected Se (6.2 ug/L) - associated samples were not detected CCB 46208/72 (tot) Sb (5.4 ug/L)- qualify associated sample detects < RL as U As (2.0 ug/L) - qualify associated sample detects < RL as U Be (0.7 ug/L) - associated samples were not detected Se (6.0 ug/L) - associated samples were not detected | PZ-205-AS Zn (tot) 100 U As (tot) 50 U As (dis) 50 U PZ-205-SS Zn (tot) 100 U PZ-206-SS, PZ-204-SS, D-87, PZ-106-SD, D-93, D-6, D-83 Zn (tot) 100 U I-68 As (tot) 50 U V (tot) 250 U PZ-106-SS As (tot) 50 U Zn (tot) 100 U
|  | CCB 47079/13 (dis) As (2.2 ug/L) - qualify associated sample detects < RL as U Se (7.7 ug/L) - qualify associated sample detects < RL as U Tl (5.1 ug/L) - associated samples were not detected CCB 47079/25 (dis) Se (8.3 ug/L) - associated samples were not detected CCB 47079/37 (dis) Se (6.4 ug/L) - associated samples were not detected CCB 47079/49 (dis) As (3.1 ug/L) - qualify associated sample detects < RL as U Se (8.2 ug/L) - associated samples were not detected | I-9 As (tot) 50 U Zn (tot) 100 U PZ-111-KS Zn (tot) 100 U As (tot) 50 U DUP05 Zn (tot) 100 U As (tot) 50 U As (dis) 50 U
CRDL Standard | CRDL results btw 70-130% 50-150% for Sb, Pb, and Tl | The CRI was in control | None
Laboratory Control Sample | LCS/LCSD: 1 per 20 project samples or each preparation batch. LCS limits within 80-120%. | LCS results within limits for total and dissolved metals | None
Lab Duplicates | RPD <20% or Absolute Diff < RL when samp/dup value < 5x RL | No lab duplicates were analyzed within this SDG. | None
Matrix Spike/Matrix Spike Duplicates | MS/MSD: 1 per 20 project samples or each preparation batch. Recoveries within lab limits. MS/MSD %RPDs <= 20%. Post Digestion Spike Recovery (when MS spike rec out of control) limits 75-125% | MS/MSD analysis for total 6010 metals performed on PZ-204A-SS Al 126%/PASS - qualify associated sample conc as J+ - Some analytes not calculated because the sample conc was > 4X the spike conc - apply qualifications to parent sample only (entire batch not qualified because another spike sample was in control) MS/MSD analysis for dissolved 6010 metals performed on PZ-204A-SS - results were in control Some analytes not calculated because the sample conc was > 4X the spike conc MS/MSD analysis for total mercury performed on PZ-201A-SS- results were in control MS/MSD analysis for dissolved mercury performed on PZ-205-AS - results were in control | PZ-204A-SS Al (tot) 4500 J+
ICP Interference Check Sample (ICS) | ICS results within 80-120%. | All criteria met. | None
Linear Range Check Standard | results within 90-110% | Linear range check standards were in control. | None
Serial Dilution | Performed on samples of a similar matrix or 1 per 20 samples. %D ≤ 10% conc ≥ 50xDL for 5-fold dilution. | Serial dilutions were performed on PZ-204A-SS, PZ-201A-SS ( tot Hg), and PZ-205AS (dis Hg). Results were in control. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP05 was collected as a field dup of I-9 - all results were in control | None
RT = Retention Time; %D = Percent Deviation; %RPD = Relative Percent Difference; %RSD = Percent Relative Standard Deviation; RRF = Relative Response Factor; CCV = Continuing Calibration Verification
TCL = Target Compound List; MS = Matrix Spike; MSD = Matrix Spike Duplicate;

## WetChem
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2044
FRACTION: | Chloride, Sulfate, Nitrate, Bromide and Iodide (EPA 300)/ Alkalinity (EPA 310.1)
MEDIA: | GROUNDWATER
SAMPLES: | PZ-103-SS, PZ-114-AS, PZ-201A-SS, PZ-204A-SS, PZ-205-AS, PZ-205-SS, PZ-206-SS, PZ-204-SS, I-68, D-87, PZ-106-SD, S-82, PZ-106-SS, I-9, D-93, PZ-111-KS, D-6, D-83, DUP05
SAMPLE DATE(S): | 4/8,9/2013
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times & Preservation | Cooler temp < 6 C. | Coolers received at 2° C. Holding time exceeded for nitrate on some samples Holding time requirements were not met for all samples. | PZ-103-SS PZ-114-AS PZ-201A-SS PZ-204A-SS PZ-205-AS PZ-205-SS PZ-206-SS PZ-204-SS NO3 R/J
Holding time = 28 days for chloride, sulfate, bromide and iodide 14 days for alkalinity 48 hrs for nitrate.
Calibration and Continuing Calibration Verification Standards | R2>0.99 CCVs 90-110% | ICVs, CCVs, and Calibrations were in control. | None
Blanks | Method blanks: 1 per 20 project samples. No analytes detected in MBs, CCBs or EBs. | Method Blank: no analytes were detected in the method blanks Equipment blank: No equipment blanks were collected for these methods. CCBs: CCB 160/47338/68 Cl 0.0438 mg/L - sample concs > 5X blank conc X DF- no qualification required. | None
Laboratory Control Sample | Recoveries within lab limits. | LCS recoveries were in control. | None
| Lab Duplicate: RPD <20% or Absolute Diff < RL when samp/dup value < 5x RL | PZ-103-SS (NO3, I, Cl, SO4, & Br) and PZ-204A-SS (Alk) were used as lab duplicates. Results were in control. | None
Lab Duplicates
Spike Sample Analysis | MS recovery within lab established limits | MS analysis for bromide, I, Cl, SO4 and NO3 was performed on PZ-103-SS and PZ-106-SD - results were in control MS analysis for alkalinity was performed on PZ-204A-SS and D-87 - results were in control | None
Low Level Control Sample | recovery within 80-120% | Performed on alkalinity only. Results were in control. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP05 was collected as a field dup of I-9 NO3 = 200% - qualify parent sample and DUP as J | DUP 05 NO3 2.1 J I-9 NO3 ND UJ

## field dups
#sys_sample_code | total_or_dissolved | chemical_name | DUP05_2013-04-09 | I-9_2013-04-09 |  | RANGE | RPD
DUP05_2013-04-09 | D | ARSENIC | 24 | 21 | 50 | 3 | 13.33%
DUP05_2013-04-09 | T | ARSENIC | 24 | 26 | 50 | 2 | 8%
DUP05_2013-04-09 | D | BARIUM | 1500 | 1400 | 250 | 100 | 6.897%
DUP05_2013-04-09 | T | BARIUM | 1400 | 1400 | 250 | 0 | 0%
DUP05_2013-04-09 | T | BROMIDE | 4 | 4 | 0.25 | 0 | 0%
DUP05_2013-04-09 | D | CALCIUM | 290000 | 270000 | 10000 | 20000 | 7.143%
DUP05_2013-04-09 | T | CALCIUM | 270000 | 270000 | 10000 | 0 | 0%
DUP05_2013-04-09 | T | IODIDE (AS I) | 0.4 | 0.4 | 1 | 0 | 0%
DUP05_2013-04-09 | D | IRON | 35000 | 34000 | 500 | 1000 | 2.899%
DUP05_2013-04-09 | T | IRON | 34000 | 34000 | 500 | 0 | 0%
DUP05_2013-04-09 | D | LEAD | 9.5 | 9 | 50 | 0.5 | 5.405%
DUP05_2013-04-09 | D | MAGNESIUM | 72000 | 71000 | 5000 | 1000 | 1.399%
DUP05_2013-04-09 | T | MAGNESIUM | 71000 | 72000 | 5000 | 1000 | 1.399%
DUP05_2013-04-09 | T | MANGANESE | 1100 | 1100 | 75 | 0 | 0%
DUP05_2013-04-09 | D | MANGANESE | 1100 | 1200 | 75 | 100 | 8.696%
DUP05_2013-04-09 | D | MERCURY | 0.064 | 0.066 | 0.2 | 0.002 | 3.077%
DUP05_2013-04-09 | T | Potassium | 20000 | 20000 | 25000 | 0 | 0%
DUP05_2013-04-09 | D | Potassium | 20000 | 20000 | 25000 | 0 | 0%
DUP05_2013-04-09 | D | SODIUM | 250000 | 240000 | 5000 | 10000 | 4.082%
DUP05_2013-04-09 | T | SODIUM | 240000 | 250000 | 5000 | 10000 | 4.082%
DUP05_2013-04-09 | T | SULFATE (AS SO4) | 0.31 | 0.26 | 0.5 | 0.05 | 17.54%
DUP05_2013-04-09 | T | VINYL CHLORIDE | 2.4 | 2.8 | 5 | 0.4 | 15.38%
DUP05_2013-04-09 | D | ZINC | 37 | 32 | 100 | 5 | 14.49%
DUP05_2013-04-09 | T | ZINC | 27 | 31 | 100 | 4 | 13.79%
DUP05_2013-04-09_DL | T | ALKALINITY, TOTAL (AS CACO3) | 1000 | 1100 | 25 | 100 | 9.524%
DUP05_2013-04-09_DL | T | NITROGEN, NITRATE (AS N) | 2.1 | 0 | 0.4 | 2.1 | 200%
DUP05_2013-04-09_DL2 | T | CHLORIDE (AS CL) | 310 | 320 | 40 | 10 | 3.175%