## VOA
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2085
FRACTION: | VOA (SW846 8260C) for the Target Compound List only
MEDIA: | GROUNDWATER
SAMPLES: | PZ-208-SS, PZ-101-SS, MW-1204, PZ-113-AS, I-73, PZ-113-AS, PZ-107-AS, PZ-116-SS, D-14, PZ-112-AS, PZ-202-SS, I-4, DUP 07 (field dup of MW-1204)
SAMPLE DATE(S): | 2013-04-12
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times, Preservation | Cooler temp < 6 ºC. Samples holding time requirement < 7 days (<14 days if HCL preserved). | Coolers received at 4° C. All holding time criteria met. | None
System Monitoring Compounds | recoveries within limits (70 - 130%) or laboratory established limits | System monitoring compound recoveries were in control. | None
Matrix Spike/Matrix Spike Duplicates | MS/MSD: 1 per 20 project samples. Recoveries and RPDs within lab limits | MS/MSD performed on PZ-208-SS all results were in control | None
Lab Control Sample/Duplicate | Recoveries within lab limits (or 70-130%). | LCS results for target compounds were within control limits | None
Blanks | Method blanks: 1 per 12 hours. No analytes detected in MB, TB, or EB. | MB 160-47057/2 all compounds were not detected MB 160-447062/2 all compounds were not detected | None
| Trip Blanks | all compounds were not detected | None
| Equipment Blanks | No field blanks were analyzed with this SDG | None
GC/MS Initial Calibration | %RSD ≤ 30%. Average RRFs > 0.050. | Initial Calibration Instrument MSL4/16/13: All %RSD criteria met. Average RRF meets requirements. | None
GC/MS Continuing Calibration | CV performed for every 12 hours per instrument. %D ≤ 25%. RRFs ≥ 0.05. | CCVIS 160-47057/3 - results were in control CCVIS 160-47062/3 - results were in control RF was greater than or equal to 0.05 ( or 0.01 for poor response compounds) for all target compounds | None
Internal Standards | IS areas of samples & blank within (-50% to + 100%). RTs < 30 seconds. | Standard recovery area within the QC limits for all three standards; and retention times were within 30 seconds of the standard for all samples that were used in this SDG. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP 07 was collected as a field dup of MW-1204 - all results were in control | None
RT = Retention Time; %D = Percent Deviation; %RPD = Relative Percent Difference; %RSD = Percent Relative Standard Deviation; RRF = Relative Response Factor; CCV = Continuing Calibration Verification
TCL = Target Compound List; TIC = Tentatively Identified Compound

## Metals
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2085
FRACTION: | Metals (SW846 6010B and 7470A/7471A)
MEDIA: | GROUNDWATER (TOTAL)
SAMPLES: | PZ-208-SS, PZ-101-SS, MW-1204, PZ-113-AS, I-73, PZ-113-AS, PZ-107-AS, PZ-116-SS, D-14, PZ-112-AS, S-53 (dissolved only), PZ-202-SS, I-4, DUP 07 (field dup of MW-1204)
SAMPLE DATE(S): | 2013-04-12
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times & Preservation | Cooler temp < 4 C. Holding Time: Hg < 28 days, all other metals < 180 days from collection. | Coolers received at 4° C. All holding time criteria met. | None
Calibration | Calibration every 24 hours Hg r^2 > 0.995 CCV every 10 samples or 2 hours ICP CV/CCV %R = 90-110% Hg CV/CCV %R = 80-120% | Hg calibrations, ICVs, and CCVs were in control Metals ICVs and CCVs were in control | None
Blanks (method blank, prep blank) | CC Blank Conc < CRQL. Method blanks: 1 per 20 project samples. No TCL detected in MB, TB, or EB. | Method blanks: MB 160-46454/1-A (total) Zn 6.6 ug/L - qualify associated sample detects < RL as U - qualify associated sample detects> RL but <5X blank conc X DF as J+ MB 160-46455/1-A (dissolved) Zn 7.1 ug/L - qualify associated sample detects < RL as U MB 160-47518/1-A: Hg (Total) - not detected MB 160-47520/1-A: Hg (dissolved) - not detected Equipment blank: No equipment blanks associated to this SDG. CCBs: CCB 47292/73 (tot) As (3.4 ug/L) - associated samples were not detected Be (0.8 ug/L) - associated samples were not detected Se (5.2 ug/L) - associated samples were not detected Tl (4.4 ug/L) - associated samples were not detected | PZ-208-SS Zn (tot) 100 U Zn (dis) 100 U Se (dis) 75 U PZ-101-SS Zn (tot) 100 U Zn (dis) 100 U As (tot) 50 U MW-1204 Zn (tot) 100 U PZ-113-AS Zn (tot) 100 U Zn (dis) 100 U I-73 As (dis) 63 J+ V (dis) 250 U PZ-113-AS Zn (dis) 100 U
|  | CCB 47292/85 (tot) As 2.0 ug/L) - qualify associated sample detects < RL as U Be (0.7 ug/L) - qualify associated sample detects < RL as U Se (3.9 ug/L) - associated samples were not detected CCB 47292/97 (tot) As 2.1 ug/L) - qualify associated sample detects < RL as U Be (0.7 ug/L)- qualify associated sample detects < RL as U Se (5.5 ug/L) - associated samples were not detected CCB 47292/109 (tot) Be (0.7 ug/L) - associated samples were not detected Se (3.3 ug/L) - associated samples were not detected | PZ-107-AS Zn (dis) 100 U As (tot) 50 U Be (tot) 25 U PZ-116-SS Zn (tot) 100 U Zn (dis) 100 U D-14 Zn (tot) 100 U Zn (dis) 100 U As (tot) 50 U PZ-112-AS Zn (tot) 100 U
|  | CCB 47506/26 (dis) As 3.3 ug/L) - qualify associated sample detects < RL as U - qualify associated sample detects> RL but <5X blank conc X DF as J+ Se (8.0 ug/L) - associated samples were not detected V (4.6 ug/L) - qualify associated sample detects < RL as U CCB 47506/38 (dis) As 2.5 ug/L) - qualify associated sample detects < RL as U - qualify associated sample detects> RL but <5X blank conc X DF as J+ Se (8.6 ug/L) - associated samples were not detected CB 47506/50 (dis) Se (6.7 ug/L) - associated samples were not detected | S-53 Zn (dis) 100 U PZ-202-SS Zn (tot) 100 J+ Zn (dis) 100 U I-4 Zn (tot) 100 U Zn (dis) 100 U As (tot) 50 U DUP 07 Zn (tot) 100 U
CRDL Standard | CRDL results btw 70-130% 50-150% for Sb, Pb, and Tl | The CRI was in control | None
Laboratory Control Sample | LCS/LCSD: 1 per 20 project samples or each preparation batch. LCS limits within 80-120%. | LCS results within limits for total and dissolved metals | None
Lab Duplicates | RPD <20% or Absolute Diff < RL when samp/dup value < 5x RL | No lab duplicates were analyzed within this SDG. | None
Matrix Spike/Matrix Spike Duplicates | MS/MSD: 1 per 20 project samples or each preparation batch. Recoveries within lab limits. MS/MSD %RPDs <= 20%. Post Digestion Spike Recovery (when MS spike rec out of control) limits 75-125% | MS/MSD analysis for total 6010 metals performed on PZ-208-SS - Some analytes not calculated because the sample conc was > 4X the spike conc MS/MSD analysis for dissolved 6010 metals performed on PZ-208-SS - results were in control Some analytes not calculated because the sample conc was > 4X the spike conc MS/MSD analysis for total mercury performed on PZ-101-SS - MS/MSD = 66%/63% - qualify associated sample as UJ- MS/MSD analysis for dissolved mercury performed on PZ-101-SS - MS/MSD = 60%/53% - qualify associated sample as J- | PZ-101-SS Hg (tot) 0.067 J- Hg (dis) 0.073 J-
ICP Interference Check Sample (ICS) | ICS results within 80-120%. | All criteria met. | None
Linear Range Check Standard | results within 90-110% | Linear range check standards were in control. | None
Serial Dilution | Performed on samples of a similar matrix or 1 per 20 samples. %D ≤ 10% conc ≥ 50xDL for 5-fold dilution. | Serial dilutions were performed on PZ-206-SS, PZ-101-SS (Hg). Results were in control. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP 07 was collected as a field dup of MW-1204 - all results were in control | None
RT = Retention Time; %D = Percent Deviation; %RPD = Relative Percent Difference; %RSD = Percent Relative Standard Deviation; RRF = Relative Response Factor; CCV = Continuing Calibration Verification
TCL = Target Compound List; MS = Matrix Spike; MSD = Matrix Spike Duplicate;

## WetChem
PROJECT NAME/NO. | West Lake Landfill Operable Unit 1
LAB: | TestAmerica - St. Louis
SDG: | 160-2085
FRACTION: | Chloride, Sulfate, Nitrate, Bromide and Iodide (EPA 300)/ Alkalinity (EPA 310.1)
MEDIA: | GROUNDWATER
SAMPLES: | PZ-208-SS, PZ-101-SS, MW-1204, PZ-113-AS, I-73, PZ-113-AS, PZ-107-AS, PZ-116-SS, D-14, PZ-112-AS, PZ-202-SS, I-4, DUP 07 (field dup of MW-1204)
SAMPLE DATE(S): | 2013-04-12
| Acceptable limits / criteria |  | Qualifiers Added?
Comments/Qualifying Actions
CRITERIA
Data Completeness, Holding Times & Preservation | Cooler temp < 6 C. | Coolers received at 4° C. Holding time requirements were met for all samples. | None
Holding time = 28 days for chloride, sulfate, bromide and iodide 14 days for alkalinity 48 hrs for nitrate.
Calibration and Continuing Calibration Verification Standards | R2>0.99 CCVs 90-110% | ICVs, CCVs, and Calibrations were in control. | None
Blanks | Method blanks: 1 per 20 project samples. No analytes detected in MBs, CCBs or EBs. | Method Blank: Alkalinity 0.250 mg/L - all sample concentrations > 5XDFXblank conc - no qualifications required Equipment blank: No equipment blanks were collected for these methods. CCBs: no analytes were detected in the CCBs | None
Laboratory Control Sample | Recoveries within lab limits. | LCS recoveries were in control. | None
| Lab Duplicate: RPD <20% or Absolute Diff < RL when samp/dup value < 5x RL | PZ-208-SS (NO3, I, Cl, SO4,Br &Alk) was used as a lab duplicate. Results were in control. | None
Lab Duplicates
Spike Sample Analysis | MS recovery within lab established limits | MS analysis for bromide, I, Cl, SO4, ALK and NO3 was performed on PZ-208-SS Br MS = 60% - qualify parent sample as J- NO3 MS = 63% - qualify parent sample as UJ- MS analysis for bromide, I, Cl, SO4 and NO3 was performed on PZ-202-SS Br MS = 132% - sample was not detected - no qualification necessary NO3 MS = 113% - sample was not detected - no qualification necessary | PZ-208-SS NO3 0.38 J- Br UJ-
Low Level Control Sample | recovery within 80-120% | Performed on alkalinity only. Results were in control. | None
Field Duplicate | Field duplicate: RPD < 35% or Absolute Diff < 5XRL when samp/dup value < 5X RL | The following field duplicates were collected with this sample set. DUP 07 was collected as a field dup of MW-1204 - NO3 was out of control at 83% - qualify associated smaples as J | all sample NO3 J/UJ

## field dups
total_or_dissolved | chemical_name | DUP 07_2013-04-12 | MW-1204_2013-04-12 |  | range | rpd
T | NITROGEN, NITRATE (AS N) | 0.24 | 0.099 | 0.02 | 0.141 | 83.19%
T | SULFATE (AS SO4) | 0.08 | 0.096 | 0.5 | 0.016 | 18.18%
T | ALKALINITY, TOTAL (AS CACO3) | 510 | 460 | 5 | 50 | 10.31%
D | BARIUM | 350 | 340 | 250 | 10 | 2.899%
T | BARIUM | 300 | 300 | 250 | 0 | 0%
T | CALCIUM | 110000 | 110000 | 5000 | 0 | 0%
D | CALCIUM | 130000 | 120000 | 5000 | 10000 | 8%
D | IRON | 5000 | 4800 | 500 | 200 | 4.082%
T | IRON | 4700 | 4800 | 500 | 100 | 2.105%
T | MAGNESIUM | 52000 | 53000 | 5000 | 1000 | 1.905%
D | MAGNESIUM | 61000 | 59000 | 5000 | 2000 | 3.333%
T | MANGANESE | 100 | 100 | 75 | 0 | 0%
D | MANGANESE | 98 | 100 | 75 | 2 | 2.02%
T | SODIUM | 12000 | 12000 | 5000 | 0 | 0%
D | SODIUM | 16000 | 15000 | 5000 | 1000 | 6.452%
T | ZINC | 32 | 33 | 100 | 1 | 3.077%
T | ACETONE | 9 | 7.7 | 20 | 1.3 | 15.57%
T | CHLOROBENZENE | 0.92 | 0 | 5 | 0.92
T | TERT-BUTYL METHYL ETHER | 1.1 | 1.1 | 5 | 0 | 0%
T | CHLORIDE (AS CL) | 17 | 15 | 4 | 2 | 12.5%